Current shut down procedure for Macromizer

Version: 1b

Date: April 08, 2004 / 16.4.2004

Authors: Hans Aerni, Vanderbilt University & Gertraud Hayn, COMET AG Switzerland

Shut down procedure (instrument will be off less than 1 hour)

· Close Cryo-valve in software (Macromizer system controller/Cryostat)

· Shut down turbos (Macromizer system controller/Vacuum/)
· Switch flight tube turbo pump off

· Switch sample inlet turbo pump off

· Wait 10 min to allow turbo pumps to spin down

· Shut down compressor (Macromizer system controller/View/ Cryostat/Turn off drive button)

· Shut down CSW-71 Helium compressor unit located in pump room (Use main switch on the front side of the instrument).

· Override security –open doors (Macromizer system controller /View/Sample Plate handling)

· Close Turbo valve (Macromizer system controller/View/ Vacuum just click in the window)

· Stop roughing pump (Macromizer system controller/Vacuum/ Rough Pump Power off button) when pressure is in the range of 10-1 mbar

· Shut down the GUI

· Shut down the server (connect via the VNC-button to the server, another windows-surface appears / open the system server window = white on black/ type Q (=quit) / close system server window / 

·  shut down the server (open the “Start” on the left bottom of the window) and the client computer .

· Turn main switch on Macromizer to off.

Shut down procedure (includes venting the instrument)

· Close Cryo-valve in software (Macromizer system controller/Cryostat)

· Shut down turbos (Macromizer system controller/Vacuum/

· Switch tube turbo pump off

· Switch chamber turbo pump off

· Wait 15 min to allow turbo pumps to spin down

· Override security –open doors (Macromizer system controller /View/Sample Plate handling)

· Close Turbo valve (Macromizer system controller/View/ Vacuum just click in the window)

· Open prechamber valve (same window as above) and pump until the pressure in the prechamber is in the 1 mbar region

· Close prechamber valve

· Open inner door

· Open nitrogen valve and flood the instrument

· Shut down CSW-71 Helium compressor unit located in pump room (Use main switch on the front side of the instrument).

· Disconnect the server, shut down the computer.

· Turn main switch on Macromizer to off

· Stop roughing pump (Macromizer system controller/Vacuum/ Rough Pump Power off button) when pressure is in the range of 10-1 mbar

· Shut down the GUI

· Shut down the server (connect via the VNC-button to the server, another windows-surface appears / open the system server window = white on black/ type Q (=quit) / close system server window / 

·  shut down the server (open the “Start” on the left bottom of the window) and the client computer .

· Turn main switch on Macromizer to off. 

Start up procedure (e.g. after power failure, instrument was off less than 1 hour)

· Turn main switch on Macromizer to on (Instrument will automatically start the rough pump followed by the turbo pumps).

· Start up the computer

· Connect to the server via VNC (button on the client computer windows surface)

· Start system server on the server (button on the server windows surface)

· Start software via system control remote on the client computer (button on the windows surface)

· Turn on CSW-71 Helium compressor unit located in pump room (Use main switch on the front side of the instrument)

· Go to Macromizer system controller

· Turn compressor on Macromizer system controller/Cryostat/Turn on drive button)

· Check temperatures (Macromizer system controller/Cryostat) If temperature in 1st shielding is around 30K and in the 2nd shielding is around 3 K

· Go to Scheduling (GUI main page/View)

· Start Regeneration task (check watch and add 5 minutes, as the time on the server is usually a little ahead)

· Start rough pump (Macromizer system controller/Vacuum/ Rough Pump Power on button)

· Open turbo valve (Macromizer system controller/View/ Vacuum just click in the window). If pressure is in the 5 mbar region

· Switch on turbo pumps ((Macromizer system controller/View/ Vacuum just click in the window)

· Wait until pressure is at least in the region of 10-5 mbar

· Let instrument to cool until detector temperature reaches 0.3 K

· Instrument is ready

If the holding time of the cryostat is reduced (due to freezing of moisture in the cryostat) from over 24h to 18hours or shorter, contact COMET Switzerland!!!

Daniel Merz

0041 31 744 9159

Gertraud Hayn

0041 31 744 9191

Flemming Skriver
0041 31 744 9145

macromizer Doc: Technical Doc

page 1

